[image: image1.png]i
\

/7 JOANLAB’



[image: image19.jpg]MG (F7)4100000391 S



[image: image20.png]


[image: image21.png]


                    [image: image30.png]



         [image: image2.jpg]



FA Series Analytical Balance
[image: image22.jpg]T~




USER MANUAL

CATALOG

ABOUT THE SCALE.....................................................................................................01
I.INSTALLATION.............................................................................................................02
 1.1 Unboxing.................................................................................................................02

 1.2 Work environment ...............................................................................02

 1.3 Installation............................................................................................03

 1.4 keyboard/Display..................................................................................03
II.OPERATION.................................................................................................................04
 2.1 Calibration..............................................................................................................04

 2.1.01Calibration weight............................................................................04

 2.1.02 Calibration Step..............................................................................04

 2.2 Function setting....................................................................................................05

 2.2.01Unit conversion ...............................................................................................05

 2.2.02 Speed adjustment..........................................................................................05

 2.2.03 Stability adjustment .......................................................................................05

 2.2.04 Counting Function .........................................................................................06

 2.2.05 Factory setting ................................................................................................07

     2.2.06 Percentage function.......................................................................................07

     2.2.07 Peak hold..........................................................................................................07

     2.2.08 Baud rate optional..........................................................................................08

III.DATA OUTPUT............................................................................................................08
IV.FAULT JUDGMENT................................................................................................10
V.MAINTENANCE...........................................................................................................11
PACKING LIST.................................................................................................................12
PACKING LIST
	NO.
	ITEM
	NUMBER

	1
	BALANCE
	1 

	2
	PAN
	1

	3
	ADPTER
	1

	4
	MANUAL
	1

	5
	CERTIFICATE OF CONFORMITY 
	1

	6
	CAL.WEIGHT
	1

	7
	PACKAGE
	1


 Note: After unpacking, please keep the balance packaging properly for later maintenance

-12-

OVERVIEW
This manual is used for FA series multifunctional analytical balance.

FA series multifunctional electronic balance is a new generation product newly developed. Using high-precision electromagnetic force balance sensor. It makes the measurement results more accurate, faster response, and fewer failures. FA series multifunctional electronic balances have:

Fast weighing. The speed can be adjusted

Easy to operate
Full range tare,clear, accumulate, overload, underload display, fault alarm within the full range;

Counting function
Optional configuration and function expansion

	Model
	FA1004
	FA1204
	FA2004
	FA2204

	Max.CAP
	100g
	120g
	200g
	220g

	Accuracy
	[image: image23.jpg]o
<
2
z
3
:
v




 

	Readability
	0.0001g (0.1mg)

	Repeat ability
	±0.0002g

	Liner 
	±0.0003g

	Stable time
	≤3s

	Operation temp
	5℃～35℃

	Pan size
	Ф80mm

	Dimensions
	300X200X295mm

	Power Supply
	Adapter

	Cal
	External（Internal optional）

	G.W
	8.0kg

	N.W
	6.0kg

	Package Size
	         425x320x415mm  

	Interface
	       RS232（USB made custom）


-01-

I.INSTALLATION
1.1. Unboxing

Please carefully open the packing box, gently take out the balance and the contents in the box, and save the packing materials for reuse.
(1) BALANCE      (2) PAN         (3)ADAPTER  
(4) MANUAL    (5)  CERTIFICATE OF CONFORMITY        (6) WEIGHT
[image: image24.png]q

Pin2 - RxD
Pin3 - TxD
Pin5 - GND




[image: image25.png]SPEEM




           1  2  3           4          5  7   6         8  9 

① Keyboard  ②Bubble  ③ Data display window  ④Pan  ⑤RS232
⑥Adjustable feet  ⑦Power outlet  ⑧Side sliding door  ⑨Top sliding door
1.2. Choice of working environment

If conditions do not permit, the working environment should be selected according to the following requirements:

(1) The studio should be kept clean and dry;

(2) The balance should be placed on a stable and fixed workbench (it is recommended to be placed on a marble platform);

(3) The workbench should be far away from doors and windows, so as to reduce the influence of airflow caused by opening windows and doors;

(4) The workbench should be set up where there is little vibration interference. The surroundings of the room are less affected by vibration, and the ideal position of the workbench should be placed;

(5) The balance should be placed in a place that is easy to cause temperature changes, 

-02-
V.MAINTENANCE

The FA series multifunctional electronic balance is a precision mechatronics intelligent measuring instrument, so it must be treated as seriously and carefully maintained as other precision instruments.

1. Do not use sharp objects (such as pencils, ball-point pens) to press keys, only use your fingers to press keys;

2. Be careful not to let objects fall from a high place on the weighing pan, so as not to damage the weighing mechanism;

3. Do not expose the balance to high humidity or dusty environment for a long time;

4. After the balance is used up, it is best to cover it with a cover to prevent dust intrusion;

5. When weighing powder and fine particles, please use it with the container to avoid dust and particles falling into the load cell under the weighing pan;

6. Please wear gloves when calibrating with weights, and do not take the weights directly with your hands;

7. Keep the balance clean and dry.

Matters needing attention when cleaning:

Before cleaning, unplug the power supply;

Do not use corrosive cleaning agents (such as solvents). Use a lint-free soft cloth dipped in water and then some neutral detergent (soap) for cleaning.

When cleaning, be careful not to let water drop into the balance;

After cleaning, wipe the balance carefully with a soft, dry, lint-free cloth.
-11-
IV.FAULT JUDGMENT
	FAULT
	REASON
	EXCLUDE

	No display
	Not connected to the power supply;

Fuse is broken;

Power transformer damage;
	Plug in the power line;

Replace the fuse;

Replacement of power transformer;

Contact factory

	Weighing unstable
	1. Bad working environment;

2. The windproof door is not closed properly;

3. There are foreign objects or scratches between the weighing pan and the work surface;

4. The power supply is unstable and exceeds the allowable value;

5. The weighing object is unstable;

(Such as absorbing moisture or evaporating)

  Plastic particles or powders are electrostatically charged

6. Static electricity generated by dry weather
	Improve the working environment and avoid the interference of vibration and airflow;

Close the windproof door;

Take out the foreign body, turn the weighing pan to prevent scratching;

External access 9V DC power adapter;

Operate with static eliminator or wristband with static electricity

	The displayed value does not match the actual weight
	The balance is not calibrated;

Not cleared before weighing;

Did not adjust the level;
	Calibrate the balance;

Press →0← to clear;

Use level feet to adjust the level.

	Err. 
	CAL ERROR
	Refer to the manual to re-operate

	Err.1
	COUNTING ERROR
	

	Err.2
	PUT PAN WRONGLY
	

	Err.3
	OVERLOAD
	


-10-

such as avoiding direct sunlight and away from heaters;

(6) Please avoid connecting with large machines or devices with interference, and avoid the interference of other devices;

(7) The dry environment is prone to static electricity, so corresponding measures can be taken to avoid the influence of static electricity on weighing.

(8) Keep the balance away from objects and equipment with magnets or capable of generating magnetic fields;

(9) The balance shall not be used in areas with explosive danger;

(10) Do not use the balance for a long time in a high humidity or high dust environment;

(11) When the balance is transferred from a colder environment to another warmer environment, the moisture in the air will be on the balance.

The internal condensation of the weighing system affects the accuracy and reliability of weighing. In order to eliminate the influence of moisture condensation, the balance can be left unplugged at room temperature for 2 hours before use.

2.3 Installation

Put the weighing pan on first, then plug the power adapter plug into the balance, then plug the power supply (input voltage 220V) and wait for it to turn on. Please place the weighing pan correctly, otherwise a prompt will appear that the weighing cannot be performed.

2.4. Buttons and display

The display and keyboard are like icons.

[image: image26.png](>>>]




(1) POWER  (2) SET  ( 3)TARE(→0/T←),CONFRIM  (4)CAL  (5)UNIT  (6)+
-03-
-   (8) Stability indicator  （9）Range progress bar
(10) Weighing  (11)Speed   (12)Stability   (13)UNIT (14) Animal weighing 

 II.OPERATION 
Before using the balance, the balance should be placed horizontally. Use the adjustment foot screw at the bottom of the balance to adjust the level, so that the bubble in the leveler is adjusted to the center of the circle, and then quickly press the POWER button to turn on the display. To ensure accurate weighing of the balance, please ensure that the balance is fully warmed up (the recommended warm-up time is one hour).

2.1 CALIBRATION
2.1.01 CAL WEIGHT(0IMLF2)
	MODEL
	CAL.WEIGHT

	FA1004
	100g

	FA1204
	100g

	FA2004
	200g

	FA2204
	200g


2.1.02 CAL.STEP
For FA series balances, please warm up for 1 hour and wait until the balance is stable before calibrating.

Take FA2004E as an example, weighing no-load,

1). Press →0← to tare first, the balance will display 0.0000g,

2). Then press the CAL button, the display shows CAL-200 flashing,

3) At this time, put a "200g" calibration weight on the weighing pan, and after a few seconds, it will display "200.0000g" calibration is complete.

Can be weighed.

-04-

3.4 The output sequence of one frame data of the balance is as follows:

	1
	Frame header(’E’)
	Fixed

	2
	Space（0X20）
	Fixed

	3
	stable state
	* (unstable),Space（stable）

	4
	Symbol
	+,-,Space

	5-13
	Weighing data
	Decimal point included,Meaningless front 0,print space

	14
	Space
	Fixed

	15-17
	Weighing unit
	

	18
	0x0d
	Fixed

	19
	0x0d
	Fixed


-09-

2.2.08 BAUD RATE CHOOSING
The steps are as below:
1.Press “SET” to enter menu,then press “SET” until it shows“BAUD”;
    [image: image3.png]



2.Press “O/T” to enter baud rate choosing,it shows“ 1200”,press “UNIT” to choose what you need,press “O/T” to confirm ON/OFF.
[image: image27.png]


III.DATA OUTPUT
3.1 SERIAL PORT CONNECTION 
BALANCE                               PC
RXD        2…………………………………… 3

TXD        3…………………………………… 2   
GND        5…………………………………… 5
3.2 The steps of connecting to computers:
Click the “start menu “
Program   

[image: image28.png]SERB.



          Attachments [image: image29.png][55]



                                         

Communication                                                                      
              Hyperterminal
2、Enter a name: ICON SELECTION,Click OK
Ignore the previous column ,Select "connection with a (N) : com 1" ,Click OK:
(1)Baud rate :1200

   (2) Data bits :8

   (3)Parity:No
   (4)Stop bit :1

   (5)Data flow control :hardware 

Click OK,display the current weighing value

3.3 Balance data output mode setting:

The output mode of the balance is continuous.

-08-
2.2.FUNCTION SETTING
2.2.01 UNIT CONVERSION
1.Press “SET” to choose;
   [image: image4.png]=z




2.Press“O/T”to enter:
   [image: image5.png]o

g




3.Press “UNIT” choose what you need,press “O/T” confirm the unit.
    [image: image6.png]


 ; [image: image7.png]-



 ; [image: image8.png]1
Iz



;  [image: image9.png]m




2.2.02 SPEED SETTING
1.Press “SET” until it shows “SPEED”

   

2.Press “UNIT” ,    highest speed,decreasing in order,     lowest speed, press                           “O/T” to confirm.
2.2.03 SENSITIVITY SETTING
1.Press “SET” until it shows“ STAB”,press “O/T” enter STAB setting.
   

2.Press “UNIT” to choose sensitivity,      Poorest Sensitivity,      High Sensitivity,press “O/T” to confirm.
-05-

2.2.04 COUNTING
There are counting functions in the balance, and you can use this for count the quantity.And please keep the products have the same weight, and the minimum weight must be ≥0.5mg.

The details operation as blow：

1.Press “SET” to enter menu,then press “SET” unit it shows“COU”;
    [image: image10.png]paca |
[





2.Press “O/T” to enter,it shows“ COU 5pcs”;
    [image: image11.png]



3.Press “UNIT” to choose sample quantity,put matched quantity,press “O/T” ,it shows：

   [image: image12.png]BES

I

- - -

=3
C
L





   4..when “COU’ disappearing,showing digital,then you can operate
    [image: image13.png]



5.After testing,press “SET” to choose“weight,pres “O/T” back to weighing.
   [image: image14.png]YECLHE ecs





-06-
2.2.05 FACTORY SETTINGS
1.Press “SET” until it shows “DEFUALT”

   [image: image15.png]dEFRULE





2.Press “O/T” to confirm.
2.2.06 PERCENTAGE FUNCTION
Press “SET” enter menu,then press “SET” until it shows“PCT”;
    [image: image16.png]



Press “O/T” enter percentage function,it shows“ PCT 100”;
  [image: image17.png]=N
=3





Put 100% sample,press “O/T” enter percentage weighing.
2.2.07 PEAK HOLD
The steps are as below:

1.Press “SET” enter menu,then press “SET” until it shows“PEAK”;
    [image: image18.png]PERH





2.Press “O/T” enter peak hold function,it shows“ ON/OFF”,press “UNIT” to choose,press “O/T” to confirm ON/OFF;
-07-
